### Second impulse response trace
In the impulse-response module SATlive supports two impulse response traces, which can be displayed at the same time.
There are two ways to assign a trace to the second trace:
- Copy the current trace to the second trace using either the F5 key or click on the Assign button in the top menu bar.
- Load the second trace from a file using the Load second IR function from the menu Tools.
If a second trace is assigned then a trace display mode selection is shown above right above the trace display. See page 2 for details.
In that selection you can select the kind of display used for the second trace.
### How to save the impulse response?
Use the Save function to write the current impulse response to a file. If enabled in Setup → Features, the info window is shown prior to saving where you can edit the trace's description prior to saving.
If you want to save the impulse response for the use with other software or for the use as complex compensation data, then you have to select the type of file from the save button's popup menu.
See page Error: Reference source not found for details about complex compensation.
### Trace display modes in the impulse – response module
There are different ways to display the impulse – response.
You can select the kind of display you want to use using a selection box located above the trace display window.
The left selection affects the current impulse – response, the right selection the latched (second) impulse – response.
The x – axis shows the time, and the y – axis shows the amplitude value in different ways.
#### Impulse response
This display shows the impulse – response over time using a linear values and a linear scaling. This is the most basic display.
The lower display always shows the full size of the impulse – response in the linear display mode.
You can use the layout setting (see page 6)to change the display of the trace.
#### Impulse – Response (log)
This kind of display shows the amplitude – values on a logarithmic scale. The logarithmic scale, a scaling close to the human reception of levels, makes it easier to detect peaks in the low level part of the impulse – response, which would be nearly invisible on a linear scale. The highest amplitude value is taken as the zero dB reference.
You can vary the display using the Display all data setting (see page 6).
#### ETC – (Energy Time Curve)
This display shows the energy content of the impulse – response. This display creates an envelope of the impulse – response, making it easier to detect global effects. The amplitude is shown using a logarithmic scale, with the highest amplitude value as reference.
You can vary the display using the Display all data setting (see page 6).
#### Schroeder Plot
The Schroeder Plot shows the backward integration of the squared impulse – response. The backward integration creates a very smooth trace, which is used as the basis for a lot of room – acoustic calculation.
A common problem with the calculation of the Schroeder Plot is noise, which will affect the result.
There are different approaches how to deal with this problem.
Use the Configure Schroeder Trace window, described on page 4 to adjust the calculation of the trace according to your needs.
##### Configure Schroeder Trace
As mentioned above (see page 3), there are some parameters which will affect the Schroeder – trace. Especial the noise floor might create problems.
There are different approaches to deal with this challenges.
SATlive offers you the possibility to affect the calculation by choosing the calculation method and the amount of averaging applied prior to the calculation of the Schroeder trace.
You can change the settings in a small window, the so called Configure Schroeder Trace window.
###### Invoke the Configure Schroeder Trace window
In the IR – module use the menu – entry Configure Schroeder Trace in the menu Tools.
In the Room – Acoustic window (see page 10) you can either use the Configure Schroeder Trace entry in the menu Room – Acoustic, or the entry with the same name in the popup menu of the upper display area. A mouse click onto the status area at the bottom of the window also shows the window.
The options at left of the window enable you to choose the desired algorithm used for the calculation of the Schroeder trace. See the page 5 for details.
The setting Averaging time on the lower left controls the amount of averaging applied to the data prior to the calculation. A longer averaging time will improve the detection of the RT60 values but it might affect the calculation of the EDT, especial when the impulse – response shows a fast decay at its beginning.
Please note that any change in this window will trigger a new calculation, which might take some time, especial in the Room – Acoustic window, where a lot of traces will be calculated.
The button Default Settings at the lower left will reset the settings to their suggested values.
###### Algorithm used for Calculation
The Configure Schroeder Trace window (see page 4) allows you to change the algorithm used for the optimization of the Schroeder trace.
- Remove noise: The mean value of the last 10% of the full impulse – response are used to calculate the mean value of the noise – floor. Afterwards this value is subtracted from the impulse response to reduce the effect of the noise.
- Use ETC: Calculate the values based on the ETC (see page 3). Otherwise the calculation will use log logarithmic impulse response.
- Limit Interval: Estimate the noise – only part at the end the measurement and excludes it from the calculation.
- Expand Trace: Extends the trace below the noise – floor using the later part of the calculated trace.
- Correct Values: Estimates the amount of valid data removed by the noise – removal and adds it during the calculation.

The image below shows some traces based on the same impulse response but calculated different settings. See the upper right area of the image for details of each trace.

#### Layout of the trace display (Display all data)
Depending on the current settings, the impulse – response trace might contain a lot of data. Therefore there are two ways to display the trace.
You can select the mode from the sub – menu Layout in the popup menu of the trace – display.
If the entry Display all data is checked, then all values will be drawn. For the full display of a large impulse – responses this might take some time due to the large amount of values.
To reduce the time needed for painting the width of the trace – display is limited to one pixel.
You can enable this option for the main trace using the entry Display all data (main trace). The entry Display all data (latched trace) in the sub – menu affects the display of the second (latched) trace.
The entry Show all data in the popup menu of a Quick – trace (see page Error: Reference source not found) allows the control of its display.
If this option is deactivated, then only one point is drawn for each pixel of the x axis. The value drawn is the value with the highest amplitude value within the time interval assigned to this pixel. If a wide range is displayed, then this display is similar to an envelope of the trace.
The display might change when you change the range displayed or if you change the size of the display.
### The Delay-Matrix
The delay matrix shows up to thirteen time values and the difference in time between each of them. You can also add a fixed delay to the values, in order compensate for a change in the reference delay after the measurement of the delay values.
#### Values used
The delay matrix always contains the current position of the ten time-markers (see page Error: Reference source not found for further details).
In the measurement of the impulse – response in the MAT module the position of the cursor, and the peak – positions in the upper and in the lower (zoom) area will show in the matrix also.
#### How to invoke
The entry delay matrix locates in the menu Marker in the pop – up menu of the trace area.
Furthermore you can find it in the marker menu in the top menu area of the delay – finder (see page Error: Reference source not found). In the MAT module you’ll find this entry in the pop – up menu of the delay section (see page Error: Reference source not found) also.
You can use the web – interface (see page Error: Reference source not found) to access the delay – matrix.
#### The Display
##### The upper Area
The upper area of the window contains a matrix that shows the difference in time and the distance calculated from the difference in time, related to the values assigned to the column and row.
The values show the delay to add to the position assigned to the row selected to reach the value assigned to the selected column.
You can select a cell by a mouse click. The values related to the selection will show in the lower area.
- To modify the label of a time – marker just right – click on its current label. This will open up a small window where you can edit the label.
- The color bar at the left indicates the polarity at the position of the marker at the moment of its assignment. A green bar indicates a value above zero and a red bar a value below zero.

##### The lower area
If a cell is selected, than the lower area shows more details about the entries assigned to the row and column.
If you select a cell in the left most column or in the topmost row, the lower area shows the information assigned to this entry. If you select an other cell, the information of both, column and row is shown in the lower area. The two upper lines contain:
- The polarity information at the moment of assignment. A green bar indicates a positive and a red bar a negative value.
- The label of the entry.
- The time value assigned to the entry
- The distance, calculated using the time value.
- The modified delay value, which is the sum of the time value and the offset value set by the time editor on the right.
The bottom line shows the difference, both in time and distance.
As long as the cursor is over the upper area the status – bar at the bottom of the window will show the row and the column of the cell at the current position of the cursor, which might differ from the cell selected.
### Room Acoustic Analysis
#### Speech Transmission Index (STI)
The Speech Transmission Index (STI) is a value used to determine the intelligibility at the position of the measurement microphone.
The value considers the loss of modulation during the transmission of the sound, but not its frequency response. SATlive uses the impulse response to calculate the STI value.
You can invoke the STI window from the Tools menu. If you’re using the menu mode complex then there will be a button STI in the upper menu bar. The loss of modulation is calculated for different frequencies used for modulation, and for different octave bands of the audio signal. There are 98 values which will be averaged using different weighting factors, based on the voice type setting, to get a single number as result.
#### The Window
The top area shows the STI value calculated. Below the STI value you’ll find the values for CSI and ALcons, both calculated based on the STI value. Right to this values the classification of the STI value according to IEC 60268.
At the upper right you can select the type of voice used for the calculation. The setting Male voice is the commonly used.
Below of this selection the part of the impulse response used for the calculation is shown. This part always starts at the start of the impulse response. Please note that for the good result this part needs to be longer as the RT60 of the room examined. You can change this setting in the STI entry of the menu. See page 11 for further information.
The center of the window shows all 98 values. Use the entry Display of the menu (see page 11) to change the type of this display.
Use the button Close at lower right to return to the impulse response module.
##### The menu bar of the STI window
The menu bar of the STI calculation offers some options for the post-process and the display of the calculated result.
On the left there is the menu STI which contains this options:
- Save trace as picture use this function to save the current display of the STI values. Select the format of the picture in the save dialog.
- Part of IR used for Calculation this setting allows you to exclude some of the tail of the impulse response from the calculation of the STI value. For a good result the length of the remaining part should exceed the RT60 of the room examined.
  The current setting will show at the display’s upper right area.
- Print use this function to print the current display, information about the impulse response used as base for the calculation and a remark, which SATlive prompts for.
- Setup Printer use this function to change the printer and its properties.
- Close closes the window and returns to the Impulse Response module.
The next entry is the submenu Display
You can select how to display the results of the STI calculation.
- Show as Matrix shows all values as numbers in a table. If the option Results colored in the menu Setup is selected then the background color of each value displays according to the value.
- Show M-Values related to Frequency Range draws the values as a function of the different audio bands used for calculation. For each frequency of modulation a single trace is drawn. You can enable the display of individual traces in the option's sub menu.
- Show M-Values related to Modulation Frequency draws the values as a function of the different frequencies used for modulation. For each band of the audio signal a single trace is drawn. You can enable the display of individual traces in the option's sub menu.
- Smoothing in the function's sub-menu you can select how much smoothing is applied to the traces.
The rightmost entry is the menu Setup.
- Results colored assign a background color to each value according to its value. This setting affects only the Show as Matrix display.
- Show grid shows a grid separating the individual values in the Show as Matrix display.
- Low Color Reference sets the threshold for the red color. Any values below this value is shown in full red in the Show as Matrix display. A value of 1.0 is shown in full green. This setting affects the layout only if the Results colored option is activated.
Hint:
- Below the display of the STI value in the top bar you can find two other values used for the judgment of intelligibility. Both are calculated based on the STI value.
  - ALcons, an abbreviation for Articulation Loss of Consonants, a number that aims on the audibility of consonants. This value shows the loss a percentage – value.
  - CIS, which means Common Intelligibility Scale, is a logarithmic conversation of the STI value. It shortens the wide range of ‘good’ STI values to allow a more exact judgment in the critical range of the intelligibility values.
#### Room Acoustic Calculation
SATlive offers a room acoustic tool to calculate room – acoustic values using an impulse response. The quality of the results relates to the quality of the impulse response used, so you should take care to create a good impulse response (use long impulse responses). The room-acoustic tool is invoked in the menu Tools (see page Error: Reference source not found).
The traces for the reverberation times are shown at the top. Below the graphic there is a table containing the numerical values of the reverberation times and other room – acoustic parameters (see page 22 for details).
The room – data manger (see page 24) on the right side of the window allows you to easily average different measurements of the same room.
The tool starts using the last selected calculation settings. See page 4 for details.
You can modify the settings using the menu entry Configure Schroeder Trace in the menu entry Room Acoustic.
##### The views of the acoustic window
The acoustic window supports two different modes.
- In the Single trace mode mode different acoustic values, all based on the same measurement (the so called master – trace), will show in the display and in the numeric display.
- Using the Multi trace mode you can display a certain room-acoustic value of more than one trace.
Use the entry Show multiple traces in the menu Display (see page 17) to toggle between the two views.
##### Managing Traces in Multi Trace Mode
Using the multi – trace mode you can display up to eleven different traces in the display.
You can control the visibility of the main – trace, of the memory – trace and of the room – data trace in their menus.
The control of the eight trace latches is similar to the quick – trace handling (see page Error: Reference source not found for details) and can be done using the popup menu.
- To assign the main – trace to a trace – latch use the entry Copy to Trace – Memory in the menu Main – Trace.
- Use the entry Load Acoustic Data to load stored acoustic data into the trace – memory.
- Show trace toggles the visibility of the trace. You can also click on the upper area of the icon to toggle the visibility of the trace.
- The entry Select this Trace brings the trace to the front and shows its values in the lower area. The name of the trace will show in the header. Click on the lower area of the icon to select the trace using the mouse.
- The entry Clear single trace will remove the acoustic data from this trace.

##### The menu Room – Acoustic
This menu is located in the menu – bar of the Acoustic Window.
- Loading will not affect Configuration select this entry to use the current settings of the Schroeder – trace. Otherwise the settings assigned to the file recalled will be used.
- Configure Schroeder Trace opens the configuration window for the calculation. See page 4 for details about this window. You can open this window also by a mouse click onto the status – bar at the bottom of the window. The status – bar shows the currently selected parameter of the calculation.
- Print prints the room acoustic parameter and the reverberation graph.
- Save Trace as Picture save the graphic to a file. You can select the kind of picture you want to save it to in the selection on the bottom of the save dialog.
- Close closes the window and returns to the impulse-response module.

##### The menu Display
Use the entries at the top this menu to select the traces to show in the display area. In the single trace mode (see page 15) you can select any combination of traces for the display. Using the multi trace mode (see page 15) allows you to select only one type of acoustic data at a time. See page 22 for a close description of the different types of acoustic data.
- Value Display open the value – display menu to configure the value – display area located in the lower half of the window.
  See page 18 for a more detailed description of this menu.
- Show multiple Traces use this setting to toggle between the single trace and the multi trace mode of the acoustic window. See page 15 for more details.
- Display Filename shows the name of the file related to the acoustic values. Otherwise the description of the file is used.
- Schroeder Frequency
  - The Schroeder Frequency is a threshold frequency based on the size of a room and on its reverberation time.
    Below of that frequency you’ll need to calculate the acoustic using single frequency waves, above the broadband approach will fit.
  - Room size enter the room size used for the calculation of the Schroeder's frequency. You can change the units used in the Setup → Features window. (see page Error: Reference source not found for details).
  - Display Schroeder Frequency marks the area below the Schroeder Frequency in the display.
##### The menu Value – Display
This menu configures the data – display in the lower half of the acoustic window.
You can access the menu either via the entry Value Display in the menu Display or by a clicking on the data – display area using the right mouse – button (popup menu).
Check the values you want to be displayed.
The height of the data – display will adapt to show all values selected.
See page 22 for a detailed description of the values available.
SATlive stores your selection.
##### The menu Range
Use the menu Range in the top menu bar to change the vertical range.
- You can choose the desired vertical display area by selecting one of the topmost entries in this menu.
- If you click on Full Range, then SATlive adjusts the display in way that all of the values are shown.
- If you enable the Autorange option, then the vertical range is adjusted to keep all of the traces displayed in the display area. This acts similar as the Full Range function, but this option adjusts the display range every time the values to display change. This setting is memorized and will be restored when you re-enter the room-acoustic tool.
- Use the Decimal Places sub-menu to adjust the display of the room-acoustic values in the lower half of the window according to your needs. Just select the desired amount of decimal places from the sub-menu.
  - No Decimal Place: The values show as whole numbers:	8 s
  - One Decimal Place: One decimal place is shown:	7.9 s
  - Two Decimal Places: Two decimal places are shown:	7.92 s
  - Three Decimal Places: The highest amount available here:	7.921 s
##### The menu Main – Trace
Use this menu to configure the display of the main trace.
- Show trace controls the visibility of the trace.
- Use Select this Trace to show the room acoustic values of this trace in the area below the display.
- Load Impulse Response loads an impulse – response from a file and converts it into acoustic data.
- Load Acoustic Data opens room – acoustic data stored in a file.
- Copy to Trace -memory will store a copy to one of the eight room data latches.
  Select the target from the sub – menu of this entry.
- Add to room data will add the acoustic values of the main – trace to the current room – data project (see page 24).
- Save Acoustic Data saves the results of the acoustic calculation based on the main – trace in a SATlive specific format.
  You can use the Memory Trace menu or the popup of the memory traces in multi trace mode (see page 15) to load and display a saved acoustic-data file.
- Export (CSV, TXT) allows you export the results either as CSV (comma separated values), which is used for spreadsheets like Excel®, or as a plain text file. You can choose the desired type of file in the lower right area of the save – dialog window.
##### The menu Memory Trace
You can load a saved result of the acoustic calculation into a second trace, which is called memory trace.
You can use the memory – trace in both views (see page 15) of the acoustic – window.
- Show Memory Trace toggles the visibility of the memory trace. Hotkey F6.
- Select this Trace (available only in multi – trace mode) brings the memory – trace to the front and shows its values in the bottom area.
- Load Acoustic Data opens a dialog where you can select the set of acoustic data which you want to use as the memory trace data.
- Duplicate Main copies the data of the current display to the memory trace.
  The shortcut for this function is F5.
- Clear empties the memory trace.
Hints:
- In the single trace mode of the acoustic window (see page 15) the values of the memory trace show as thin lines and its values are displayed using default font, while the values of the main impulse response will show as thinker lines and the values will displayed using a bold font.
- If valid acoustic data has been loaded then the file name is shown in the lower left area of the acoustic-calculation window.
##### The menu room – data
This menu controls the room – data manager. See page 24 for details about the room – data manager.
- Clear room data: Clear all data in the current room – data file. See page 24 for details for the room – data file.
- Show room data graphs: Select this entry to show the averaged values of the traces added to the current room – data set. The display will show the averaged values and their standard deviation is indicated by vertical bars.
- Check Quality: If checked SATlive will verify the quality of the values and will only add values with good quality to the room data file.
- Set description: This entry opens an editor window where you can edit the description of the current room – data set.
- Load room data: Use this entry to load a previously saved room – data set from a file.
- Save room data: This entry saves the current room – data set using the filename used during the previous loading or saving. It is enabled only if the current room – data set has either been saved or loaded before.
- Save room data as: Use this entry to save the current room – data set to a file. This entry opens a file – selection window where you can select the folder and name to store the data to.
- CSV Export: Use this entry to export the current room – data set as CSV data. You can import CSV data easily into spreadsheets, like Excel ©.
  After the selection of the folder and name for saving the file to, a window will show up where you can edit the parameters and the values used for the export. Note your settings, because the spreadsheet – software might ask for it.
- Show Room Project Area: Use this entry to show or hide the Room – Project (see page 24) area. The area will show up automatically when you open a room – project.
##### Room – Acoustic Parameters
The numeric results of the room acoustic calculation are shown in a table below the graphic display of the reverb – time values. The values are calculated for the whole frequency range and for octave-band filtered signals.
You can use the menu Value Display (see page 18) to hide the entries you’re not interested in. The height of the area will adapt to your selection.
- RT60 or reverberation time is the most common value used in room acoustics.
  It shows the calculated time that a sound would need to decay to 60dB below it's initial level. This time is calculated based on a certain range of the decay trace, depending on the signal to noise ratio available.
  - RT60(30) is based on the slope between -5dB and -35dB.
  - RT60(20) is based on the slope between -5dB and -25dB.
  - RT60(10) is based on the slope between -5dB and -15dB.
  - Optimum values for the RT60 depend on the size of the room. The literature suggests approx 2 seconds for music and approx 1 second for speech.
  - The correlation of the result, is used to determine the quality of the calculation.
    A value shown in italic font indicates a low correlation. You can display the correlation of each value by placing the cursor on an entry and holding the left mouse button down.
- EDT means Early Decay Time, which is another reverberation time. This time is calculated using the range from 0dB to -10dB. Under normal circumstances the EDT is the reverberation time that you 'feel' in a room, while the RT60 is a more technical value used for calculations. For music the EDT value should be approx 2.2 seconds.
- C50 and D50 are values which relate the energy arriving in the first 50 milliseconds compared to the whole energy (D50) or compared to the energy arriving after the first 50 milliseconds (C50). This values are useful hints for intelligibility. Values greater then 0dB (C50) or 50% (D50) indicate a good intelligibility.
- The calculation of the C80 and D80 values is similar to the calculation of the C50 and D50 values, with the difference that they relate to the energy arriving in the first 80 milliseconds. This values are used for judging the musical sound of a room. A good music reproduction is archived if the C80 value is between -3dB and +1dB, which is the same as a D80 value of around 50%.
- The center time marks the center of gravity of the energy-time area. It is a measure for position of the main energy in time. Its value should be below 80 milliseconds for speech and between 100 and 150 milliseconds for music reproduction.
- The Schroeder's frequency depends on the room size, which can be set in the display menu (see page 17) of the room-acoustic window. You need to take care of discrete room – modes only at frequencies below the Schroeder's frequency.
  The Schroeder's frequency is shown in the status – bar at the bottom of the room-acoustic window.
- The Bass – Ratio (BR) shows the relation of the sum of the RT60(20) values for the 125Hz and 250Hz band related to the sum of the values for 500Hz and 1kHz.
  This value is an indicator how “boomy” the room will sound. Values above one will indicate intelligibility problems.
Hint:
- You can use the Decimal Places sub-menu in the range menu (see page 18) to adjust the display of the values.
- You can display the regression – lines for the full – range Schroeder – plot using the Configure Schroeder Trace window (see page 4).
##### The room – data manager
This tool is attached to the room – acoustic window (see page 14).
It eases the averaging of different measurements of the same location. It supports saving and loading of room – data sets as well as export and display of the results.
The tool averages the four reverb – time values (RT6010, RT6020, RT6030 and ETC).
The so called r – value, which is an indicator for the quality of the approximation used for the calculation of the values can be used as a threshold for the use of a specific value.
The visibility of this area is controlled by the Show Room Project Area setting in the menu Room Data Manager (see page 21). The area will show up automatically when you load a room – project.
###### Adding data
To add the values of a measurement to the room – data set, load the impulse – response into the room – acoustic window using the load command located in the menu room – acoustic (see page 16 for details about this menu). After that, add it to the room – data set using the entry Add to room data in the menu room – acoustic (see page 16).
If the data you want to add match a data – set still in the room – data set then SATlive will show a confirmation window, otherwise the data will be added directly.
###### Display of the data
The averaged values are shown at the top of the right area. SATlive shows the description of the impulse – responses used for the averaging. Depending on the Check Quality setting in the menu Room – Data Manager (see page 21) values are shown only if the quality indicator (‘r – value’) of the values is over the quality – criteria.

If the option Show room data graphs is enabled in the menu Room – Data manager (see page 21) then the calculated traces will show in the display area on the upper left, with the standard deviation shown as a vertical bar.
The CSV file or the text file generated by the room – data manager contains all numerical values.